## Slide 1
Cover inlet to keep out insects
Disconnect here to allow 2537
To make zero air by vacuum for calibration
Cap unused line
Disconnect here and cap ¼ PFA tubing
Use new, clean ¼ PFA or FEP tubing to connect to the filter
Then direct the tubing down to the bottom right corner
Make a small hole in the air vent screen and have the tubing
Extend out the side of the 1130 box
Disconnect cable
To analyzer

## Slide 2
For the 1130 Pump Module
Pump Power Off
All other switches in auto – to keep outside modules warm and dry
Except – set external heater to OFF
Disconnect the cable to the analyzer
Keep power on to the 1130 and 1135 outdoor units to keep heaters on and heated line on
Set to off
Set to off

## Slide 3
For the 2537
Must turn auto-calibration from off to ON
Set Calibration interval to 36 hours
Set next calibration to a future time